## UART-AT
Command name | Direction | Description | Value type | DEC/HEX | Value range | Default | Persistent memory
STATUS
AT | r | NOP | - | - | - | - | no
ATVERSW | r | Return the current software version number | STRING | DEZ | <MAJOR.MINOR.PATCH> | - | no
ATVERHW | r | Returns the system hardware as a HEX value of the form PRDTx where P=PartID and R=ChipRevision and DT= DeviceType | UINT16 | HEX | <0xPRDT> PR = 40 DT = 15 | 0x4015 | no
ATEVENT | r | Read the 16-bit EVENT register, clearing the entire register after the read | UINT16 | HEX | Bit0 = OCC: An occupancy (OCC) event has occurred since the last register read Bit1 = DIM: An adjustment to the dimming level (DIM) has occurred since the last register read Bit2 = LED23M: An adjustment to the color tune level (LED23M) has occurred since the last register read Bit3 = not used Bit4 = SCENE: A scene has triggered since the last register read Bit5 = DL_LIMIT: the lumen controller has reached a limit, target value can not reached Bit6 = CT_LIMIT: the color temperature controller has reached a limit, target value can not reached | 0x0000 | no
ATTEMP | r | Read the current device temperature in degrees Celsius | SINT8 | DEZ | - | - | no
ATXYZC | r | Read calibrated X, Y, and Z data | FLOAT ARRAY | DEZ | <XXX.X, YYY.Y, ZZZ.Z> | - | no
ATSMALLXYC | r | Read calibrated x and y for CIE 1931 color gamut | FLOAT ARRAY | DEZ | <xxxxx.xxxx, yyyyy.yyyy> | - | no
ATUVPRIMEC | r | Read calibrated u’, v’ and u, v for CIE 1976 color gamut | FLOAT ARRAY | DEZ | <u'u'u'u'u'.u'u'u'u',v'v'v'v'v'.v'v'v'v', uuuuu.uuuu,vvvvv.vvvv> | - | no
ATDATA | r | read all six raw values: red, green, blue, ir, dark, clear | INT ARRAY | DEZ | <r, g, b, ir, d, c> | - | no
ATDUVC | r | Read delta uv values | FLOAT | DEZ | XXXXX.XXXX | - | no
ATESP | w/r | Read the single 16 bit sum of ESP1 board devices available. If you set the corresponding bit to high, it will disable the internal state machine and stops communication. Write zero starts a reinitialization. | UINT16 | HEX | Bit0 = TSL4531 Bit1 = ENS210 Bit2 = HTU20D Bit3 = CCS811 | - | no
ATTEMPX | r | Read the external temperature in degrees Celsius | FLOAT | DEZ | -40 - 125°C (format: XXX.XX) | - | no
ATRHX | r | Read the external humidity in percent | FLOAT | DEZ | 0 - 100% (format: XXX.XX) | - | no
ATECO2 | r | Reads the eCO2 [in ppm] value from the external gas sensor, if available | UINT16 | DEZ | 400ppm - 64000ppm 0 - means, that values will be measured, but internal calibration is not finished (wait 1min) | - | no
ATTVOC | r | Reads the TVOC [in ppb] value from the external gas sensor, if available | UINT16 | DEZ | 0ppb - 64000ppb | - | no
PWM
ATLIGHT | w/r | enables/disables the light (pwm output) | UINT8 | DEZ | 1 = ON, 0 = OFF | 1 | yes
ATPWMFREQ | w/r | Set/Get the PWM frequency in Hertz, SYNC-Input is only on 886Hz enabled | UINT16 | DEZ | 100 - 4000Hz | 886 | yes
ATPWMx | w/r | Set PWMx Duty Cycle target as an integer percentage value between 0 and 100, removing it from any control loop. Returns always the current configured pwm value in percent (x = 1 .. 3) | UINT8 | DEZ | 0 - 100 102 - release pwm back into loop control | PWM1 = 0 PWM2 = 50 PWM3 = 50 | yes
ATXPWMx | w/r | Set PWMx Duty Cycle target as an integer percentage value between 0 and 100, removing it from any control loop. Returns always the current configured pwm value in percent. Note: The difference to the command ATPWMx is that here the DIM value is multiplied. (x = 1 .. 3) | UINT8 | DEZ | 0 - 100 102 - release pwm back into loop control | PWM1 = 0 PWM2 = 50 PWM3 = 50 | yes
ATPWMOVR | r | Read the single 16 bit sum of PWMs to identify which have been over-ridden and removed from any closed-loop control loop | UINT8 | HEX | Bit0 = ATXPWM1 active Bit1 = ATXPWM2 active Bit2 = ATXPWM3 active Bit4 = ATPWM1 active Bit5 = ATPWM2 active Bit6 = ATPWM3 active | 0x00 | no
ATDIM | w/r | set dim level in percent (0% off, 100% full brightness) | UINT8 | DEZ | 0 - 100 | 0 | yes
ATLED23M | w/r | Set the value for PWM2. PWM3 = 100 - PWM2. The sum is 100 | UINT8 | DEZ | PWM2: 0 - 100 | 50 | yes
ATRAMP | w/r | Set ramp time for lighting adjustments in percent. Time configuration is used for ATXPWM, ATDIM, ATLED23M, occupancy and Light on/ff | UINT8 | DEZ | 0: no ramp used 1 - 100: ramp time in percent (1s - 10s) example: DIM 50% - 0% with ATRAMP=50 takes 2.5s | 50 | yes
CONTROL
ATINTTIME | w/r | Set sensor integration time. Integration time = <value> x ~2.8msecs. | UINT8 | DEZ | 1-255 | 20 | yes
ATGAIN | w/r | Set sensor gain: 0=1X gain, 1=3.7X, 2=16X, 3=64X | UINT8 | DEZ | 0-3 | 1 | yes
ATLED0 | w/r | Enables or disables the indication led | UINT8 | DEZ | 0 - LED off / 1 - LED on | 1 | yes
ATLAI | w/r | Read lower light address identifier | UINT16 | DEZ | 0 - 65535 | 65535 | yes
ATLAIE | w/r | Read higher light address identifier | UINT16 | DEZ | 0 - 65535 | 65535 | yes
ATFRST | w | Factory Reset. Stored values are reset to ‘Factory’ defaults. Afterwards a software reset is started. | - | - | - | - | no
ATSRST | w | Software reset | - | - | - | - | no
ATBLINK | w/r | Starts the light on a low-high-low output blink to allow identification during a commissioning or maintenance process | UINT8 | DEZ | 1 = ON, 0 = OFF | 0 | no
ATPERSMEM | w/r | Enable/Disable writing to the persistent memory | UINT8 | DEZ | 0 - disabled, 1 - enabled | 0 | yes
CORRELATED COLOR TEMPERATURE (CCT)
ATCT | w/r | Enables/disables color tuning | UINT8 | DEZ | 1 = ON, 0 = OFF | 0 | yes
ATCCTT | w/r | Set the color control target value in integer (in Kelvin) | UINT16 | DEZ | 400 - 15000 | 2700 | yes
ATCCTC | r | Return the calibrated CCT value | UINT16 | DEZ | 400 - 15000 | - | no
DAYLIGHT HARVESTING / ILLUMINATION CONTROL
ATDL | w/r | Enables/disables daylight harvesting | UINT8 | DEZ | 1 = ON, 0 = OFF | 0 | yes
ATLUXT | w/r | Set illumination target LUX value | UINT16 | DEZ | 0 - 64000 | 400 | yes
ATLUXS | w/r | Switch between lux values from internal sensor and external ambient light sensor | UINT8 | DEZ | 0 - auto detect (higher priority is external sensor) 1 - internal sensor 2 - external ambient light sensor | 0 | yes
ATLUXC | r | Reads internal or external LUX value in dependency of ATLUXS Information: External ALS LUX value will be calculated every 700ms, Internal ALS LUX calculation depence on configured integration time | UINT16 | DEZ | 0 - 64000, 65535 - LUX value is too high | - | no
ATDMLTL | w/r | Configures the lowest supported DIM output. It will be used for normal DIM command and DL control. So we can support external dimmers which supports not the complete range from 0 - 100% (dimming 10 - 100%) | UINT8 | DEZ | 0 - 50 | 0 | yes
DIM AND AUX
ATDE | w/r | Enable/disable DIM pin (0_10V_DIM) | UINT8 | DEZ | 1 = ON, 0 = OFF | 0 | yes
ATDMD | w/r | Set DIM pin mode. Only used if ATDE enabled! | UINT8 | DEZ | 0: DIM mode: range from 0% to 100%, direct regulation of the brightness of the pwm outputs 1: LUX mode: range from 0% to 100%, direct regulation of the illumination [LUX]. Ranges can be set via ATDMINLUX and ATDMAXLUX. Overwrites the ATLUXT command | 0 | yes
ATDMIN | w/r | ATDMIN sets low end of 0-10V DIM pin, using an integer percentage of 10V. For example ATDMIN=20 sets the lower DIM pin (0_10V_DIM) range to 2.0V (10V x 20%) for 0% brightness. | UINT8 | DEZ | 0 - ATDMAX | 0 | yes
ATDMAX | w/r | ATDMAX sets high end of 0-10V DIM pin, using an integer percentage of 10V. For example ATDMAX=80 sets the higher DIM pin (0_10V_DIM) range to 8.0V (10V x 80%) for 100% brightness. | UINT8 | DEZ | ATDMIN - 100 | 100 | yes
ATDMINLUX | w/r | Set the current minLUX target value for the DIM pin range. | UINT16 | DEZ | 0 - ATDMAXLUX | 0 | yes
ATDMAXLUX | w/r | Set the current max LUX target value for the DIM pin range. | UINT16 | DEZ | ATDMINLUX - 64000 | 64000 | yes
ATPCTLUX | w/r | Set the lux target value in percentage between ATDMINLUX and ATDMAXLUX | UINT8 | DEZ | 0 - 100, can read back also 101, which means that the internal set target value is out of range | 0 | no
ATDVAL | r | Reads the current analog value of the DIM pin in digits | UINT16 | DEZ | 0 - 1023 | - | no
ATXE | w/r | Set AUX input function enable. When set to 0, all ATXMD operations are held in default mode. | UINT8 | DEZ | 1 = ON, 0 = OFF | 0 | yes
ATXMD | w/r | Set AUX pin mode. When ATXE=0, all ATXMD operations are disabled. | UINT8 | DEZ | 0: AUX mode: This is for AUX pin usage as a 0-3V input that can scale the DIM pin input. 1: OCC mode: This is for the AUX pin to act as a digital occupancy input (50% threshold) for the OCC timer. 1= occupancy, 0=no occupancy. See also ATOCCT 2: CT mode: This is for AUX pin usage as CCT target. Overwrites the ATCCTT command 3: CT_SEL mode: Selected CCT target values will be configured in 300mV steps width | 0 | yes
ATXMIN | w/r | Sets lower end of 0-3V AUX pin, using an integer percentage of 3V. For example if ATXMIN=25 sets an AUX pin range from 0.75V (25% x 3v) to 3V for 0-100% scaling. | UINT8 | DEZ | 0 - ATXMAX | 0 | yes
ATXMAX | w/r | Sets higher end of 0-3V AUX pin, using an integer percentage of 3V. For example if ATXMAX=75 sets an AUX pin range to 2.25V (25% x 3v) to 3V for 0-100% scaling. | UINT8 | DEZ | ATXMIN - 100 | 100 | yes
ATXMINCCT | w/r | Sets min CCT for minimum AUX value in Kelvin | UINT16 | DEZ | 400 - ATXMAXCCT | 400 | yes
ATXMAXCCT | w/r | Sets max CCTT for AUX pin=3V in Kelvin | UINT16 | DEZ | ATXMINCCT -15000 | 15000 | yes
ATXVAL | r | Reads the current analog value of the AUX pin in digits | UINT16 | DEZ | 0 - 1023 | - | no
ATOCCT | w/r | Set Occupancy timeout value for ATXMD = 1 | UINT8 | DEZ | 1 - 10 [minutes] | 10 | yes
ATLOG | w/r | Set alternate dimming curve (DALI approximation) | UINT8 | DEZ | 1 = Logarithmic 0 = Linear | 0 | yes
CALIBRATION VALUES
ATNORMGAIN | w/r | Set/Get the gain which the calibration values were measured | UINT8 | DEZ | same as ATGAIN | 1 | yes
ATNORMINTT | w/r | Set/Get the integration time which the calibration values were measured | UINT8 | DEZ | same as ATINTTIME | 59 | yes
ATIRXS | w/r | Write IR scalar for value X | FLOAT | DEZ | - | p2ram value 0.0 | yes
ATIRYS | w/r | Write IR scalar for value Y | FLOAT | DEZ | - | p2ram value 0.0 | yes
ATIRZS | w/r | Write IR scalar for value Z | FLOAT | DEZ | - | p2ram value 0.0 | yes
ATPMxy | w/r | Write 3x3 production matrix to flash, x,y = [0..2] | FLOAT | DEZ | - | p2ram value, 1, 0, 0 0, 1, 0 0, 0, 1 | yes
ATAMxy | w/r | Write 3x3 application matrix to flash, x,y = [0..2] | FLOAT | DEZ | - | 1, 0, 0 0, 1, 0 0, 0, 1 | yes
ATCMxy | r | Read 3x3 calibration matrix (AM * PM) from RAM, x,y = [0..2] | FLOAT | DEZ | - | - | no
ATLXSL0 | w/r | Reads and writes a calibration scalar K0 for external ambient light sensor (TSL2572 only), Formula LUX = K0 * ADC0 - (K1 * ADC1) | FLOAT | DEZ | ≥ 0 | 0.2178 | yes
ATLXSL1 | w/r | Reads and writes a calibration scalar K1 for external ambient light sensor (TSL2572 only) Formula K0 > 0: LUX = K0 * ADC0 - (K1 * ADC1) Formula K0 = 0: LUX = K1 * ADC1 | FLOAT | DEZ | ≥ 0 | 0 | yes
FIRMWARE UPDATE
ATFWU | w | Starts firmware update process and transfer the bin file checksum | - | - | - | - | no
ATFW | w | Download new firmware. Up to 10 bytes of FW image at a time (20 hex bytes with no leading or trailing 0x). Repeat command till all 56Kbytes of firmware are downloaded. | - | - | HEX STRING (without 0x), max 10 bytes | - | no
ATFWS | w | Tests the checksum on the non-active FW partition and, if correct, switches active partition. This is a toggle and can be used to toggle between the 2 FW partitions. Note: the first 5 bytes in page 0 are not touched. It is only a temporary switch and must be used to check the new firmware whether the communication works! | - | - | - | - | no
ATFWL | w | This command locks the current firmware to starts on power cycles. It rewrites the first five bytes in page0! | - | - | - | - | no
ATFWC | r | This command gives information about the current firmware state | UINT8 | HEX | Bit0 - checksum of non-active firmware OK Bit1 - error occured Bit2 - is bank 1 active Bit3 - not used Bit4 - current firmware is locked Bit5 - 56kBytes transferred Bit6 - not used Bit7 - Firmware update active | - | no
ATFWA | w | Only for backward compatibility to support old firmware update mechanism. Always returns with OK. Because of flash devices it is not possible to increment the address separatly (Page erase necessary!) | - | - | - | - | no
SCENE
ATTIMENOW | w/r | Set the current time with 1 minute resolution | UINT16 | DEZ | 0 - 1439, 65535 - disabled | 65535 | no
ATDOW | w/r | Set day of week | UINT8 | DEZ | 0 (Monday) - 6 (Sunday) 255 - disabled | 255 | no
ATCCLR | w | Restores scene command defaults in the specified scene. | UINT16 | HEX | one bit for each scene. 1 = reset | - | no
ATSnDAY | w/r | Set scene n day(s) of week for scene operation, as a decimal sum | UINT8 | HEX | a 1 in a bit location activates that day. lsb=1=Monday operation, etc. (e.g. 07 hex=Wed,Tue, Mon) 0x00 - days not configured | 0x00 | no
ATSnTIME | w/r | Set scene n start time | UIN16 | DEZ | 0 - 1439: 1 minute steps from Midnight 65535 - scene time not configured | 65535 | no
ATSnCCT | w/r | Calls the ATCCTT command at scene trigger time | UINT16 | DEZ | see command ATCCTT, 65535 - value not used | 65535 | no
ATSnLUX | w/r | Calls the ATLUXT command at scene trigger time | UINT16 | DEZ | see command ATLUXT, 65535 - value not used | 65535 | no
ATSnDIM | w/r | Calls the ATDIM command at scene trigger time | UINT8 | DEZ | 0 - 100, 255 - value not used | 255 | no
ATSnLED23M | w/r | Calls the ATLED23M command at scene trigger time | UINT8 | DEZ | 0 - 100, 255 - value not used | 255 | no
ATSnPWMx | w/r | Calls the ATXPWMx command at scene trigger time | UINT8 | DEZ | 0 - 100, 255 - value not used | 255 | no
ATSnRAMP | w/r | Sets ramp duration time in minutes for the commands: - ATSnPWMx, ATSnDIM, ATLED23 - ramp will be handled by RAMP-generator - ATSnLUX, ATSnCCT - ramp is handled by scene loop (every second) and can be stopped by configuring of new target values This command acts only for the scene trigger time and not global like the ATRAMP command | UINT8 | DEZ | 0 - 240 255 - value not used | 255 | no
ATSnFLAGS | w/r | This command is used to enable/disable configuration flags like daylight control, color tuning, dimming, auxiliary and calls directly the related commands: ATDL, ATCT, ATDE, ATXE | UINT16 | HEX | 0x0001 – Enable DL 0x0002 – Enable CT 0x0004 – Enable DIM pin 0x0008 – Enable AUX pin functions 0xFFFF - value not used | 0xFFFF | no
ATSnON | w | Set scene n to be on/off. | UINT8 | DEZ | 1 = ON, 0 = OFF | 0 | no
ATTSON | r | Read scenes on/off status. | UINT16 | HEX | each bit for one scene | 0x0000 | no
UART - I²C bridge
ATI2C | w | Perform a single write to the I2C bus at the specified 7-bit device using the 8-bits of specified data. Supports also a single register set (without value) | UINT8 ARRAY | - | ATI2C=<7-bit_address>, <8-bit_reg>,[<8-bit_data>] - [<8-bit_data>] is optional examples: - ATI2C=50,0x12, 0xFF: writes value 0xFF to register 0x12 on device address 50 - ATI2C=50,0x12,: set register 0x12 on device address 50 |  | no
ATI2C | r | Perform a single read from the I2C bus at the specified 7-bit device address and 8-bit register, returning 8-bits of data or ERROR | UINT ARRAY | HEX | ATI2C=<7-bit_address>, <8-bit_reg> |  | no
ATI2CB | w | Starts a burst transfer up to 10 bytes for writing and reading. At first the data will be written and after then read sequence will be started | UINT ARRAY | HEX | ATI2CB=<7-bit_address>,<read_length>[,<write_data>] - first two parameter can be hex or dez, - <write_data> must be hex without leading 0x, but is optional examples: ATI2CB=0x49,2: reads two byte from addr 0x49 ATI2CB=50,4,5BF3: writes 0x5B and 0xF3 to address 50 and read 4 bytes after this |  | no